{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["废水比对"],"not_rendered":{"external_links":1,"defined_names":1},"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"废水比对","index":0,"n_rows":166,"n_cols":10,"rows_kept":166,"merged":[[1,0,17,0],[1,1,17,1],[18,0,33,0],[18,1,33,1],[34,0,50,0],[34,1,50,1],[51,0,67,0],[51,1,67,1],[68,0,71,0],[68,1,71,1],[72,0,84,0],[72,1,84,1],[85,0,101,0],[85,1,101,1],[102,0,118,0],[102,1,118,1],[119,0,122,0],[119,1,122,1],[123,0,131,0],[123,1,131,1],[132,0,148,0],[132,1,148,1],[149,0,165,0],[149,1,165,1]],"flags":["formulas_cached_only"]}
["任务名称","企业名称","监测点","监测项目","监测日期","样品类型","手动监测值","自动监测值","比对结果","技术要求"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总磷","2023-12-04","实际样","0.2","0.215","0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总磷","2023-12-04","实际样","0.2","0.225","0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总磷","2023-12-04","实际样","0.2","0.230","0.03mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总磷","2023-12-04","标准样","0.5","0.482","-3.6%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","化学需氧量","2023-12-04","实际样","25","25.9","0.9mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","化学需氧量","2023-12-04","实际样","25","24.3","-0.7mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","化学需氧量","2023-12-04","实际样","25","24.5","-0.5mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","化学需氧量","2023-12-04","标准样","50","50.6","1.2%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","氨氮","2023-12-04","实际样","1.5","1.441","-0.06mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","氨氮","2023-12-04","实际样","1.5","1.397","-0.1mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","氨氮","2023-12-04","实际样","1.5","1.445","-0.05mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","氨氮","2023-12-04","标准样","5","4.817","-3.66%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总氮","2023-12-04","实际样","7.86","6.563","-16.5%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总氮","2023-12-04","实际样","8.05","6.432","-20.1%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总氮","2023-12-04","实际样","7.80","5.991","-23.19%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","总氮","2023-12-04","标准样","15","13.922","-7.19%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-014苏州市相城区黄埭污水处理有限公司","苏州市相城区黄埭污水处理有限公司","废水排放口(S1)","pH值","2023-12-04","实际样","7.0","6.57","-0.43pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总磷","2023-11-27","实际样","0.2","0.186","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总磷","2023-11-27","实际样","0.2","0.195","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总磷","2023-11-27","实际样","0.2","0.194","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总磷","2023-11-27","标准样","0.5","0.483","-3.4%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总氮","2022-11-27","实际样","7.22","6.810","-5.68%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总氮","2023-11-27","实际样","7.57","6.720","-11.23%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总氮","2023-11-27","实际样","7.72","7.929","2.71%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","总氮","2023-11-27","标准样","15","16.269","8.46%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","氨氮","2023-11-27","实际样","1.5","1.669","0.17mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","氨氮","2023-11-27","实际样","1.5","1.570","0.07mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","氨氮","2023-11-27","实际样","1.5","1.574","0.07mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","氨氮","2023-11-27","标准样","5","4.820","-3.6%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","化学需氧量","2023-11-27","实际样","25","25.8","0.8mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","化学需氧量","2023-11-27","实际样","25","25.0","0mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","化学需氧量","2023-11-27","实际样","25","25.0","0mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-007苏州市相润排水管理有限公司-城西污水厂","苏州市相润排水管理有限公司","废水排放口(S1)","化学需氧量","2023-11-27","标准样","50","53.4","6.8%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总磷","2023-11-21","实际样","0.2","0.220","0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总磷","2023-11-21","实际样","0.2","0.227","0.03mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总磷","2023-11-21","实际样","0.2","0.219","0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总磷","2023-11-21","标准样","0.5","0.502","0.4%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总氮","2023-11-21","实际样","11.9","12.339","3.69%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总氮","2023-11-21","实际样","12.2","11.536","-5.44%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总氮","2023-11-21","实际样","11.7","11.884","1.57%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总氮","2023-11-21","标准样","15","15.476","3.17%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","氨氮","2023-11-21","实际样","1.5","1.377","-0.12mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","氨氮","2023-11-21","实际样","1.5","1.444","-0.06mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","氨氮","2023-11-21","实际样","1.5","1.531","0.03mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","氨氮","2023-11-21","标准样","5.0","5.157","3.14%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","化学需氧量","2023-11-21","实际样","25","25.66","0.66mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","化学需氧量","2023-11-21","实际样","25","24.82","-0.18mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","化学需氧量","2023-11-21","实际样","25","24.65","-0.35mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","化学需氧量","2023-11-21","标准样","50","51.48","2.96%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-004苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","pH值","2023-11-21","实际样","7.1","7.03","-0.07pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总磷","2023-09-22","实际样","0.2","0.198","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总磷","2023-09-22","实际样","0.2","0.179","-0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总磷","2023-09-22","实际样","0.2","0.183","-0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总磷","2023-09-22","标准样","1","0.983","-1.7%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","pH值","2023-09-22","实际样","7.25","6.9","-0.35pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","氨氮","2023-09-22","实际样","1.5","1.67","0.17mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","氨氮","2023-09-22","实际样","1.5","1.67","0.17mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","氨氮","2023-09-22","实际样","1.5","1.65","0.15mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","氨氮","2023-09-22","标准样","10","9.26","-7.4%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","化学需氧量","2023-09-22","实际样","25","22.49","-2.51mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","化学需氧量","2023-09-22","实际样","25","22.6","-2.4mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","化学需氧量","2023-09-22","实际样","25","23.48","-1.52mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","化学需氧量","2023-09-22","标准样","100","100.9","0.9%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总氮","2023-09-22","实际样","4.83","5.76","19.25%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总氮","2023-09-22","实际样","4.85","5.76","18.76%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总氮","2023-09-22","实际样","4.77","5.992","25.62%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-045苏州市相润排水管理有限公司-城区污水厂","苏州市相润排水管理有限公司（城区污水处理厂）","总排口","总氮","2023-09-22","标准样","20","17.723","-11.39%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-043苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","氨氮（NH3-N）","2023-09-07","实际样","1.5","1.429","-0.07mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-043苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","氨氮（NH3-N）","2023-09-07","实际样","1.5","1.427","-0.07mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-043苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","氨氮（NH3-N）","2023-09-07","实际样","1.5","1.448","-0.05mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-043苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","氨氮（NH3-N）","2023-09-07","标准样","5","4.886","-2.28%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总氮（以N计）","2023-05-26","实际样","6.78","7.537","11.17%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总氮（以N计）","2023-05-26","实际样","6.93","7.782","12.29%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总氮（以N计）","2023-05-26","实际样","6.85","7.351","7.31%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总氮（以N计）","2023-05-26","标准样","10","10.352","3.52%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-05-26","实际样","0.2","0.190","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-05-26","实际样","0.2","0.192","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-05-26","实际样","0.2","0.194","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-05-26","标准样","0.5","0.489","-2.2%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","pH值","2023-05-26","实际样","7.1","6.72","-0.38pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","化学需氧量","2023-05-26","实际样","25","25.582","0.58mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","化学需氧量","2023-05-26","实际样","25","24.552","-0.45mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","化学需氧量","2023-05-26","实际样","25","24.909","-0.09mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-028苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","化学需氧量","2023-05-26","标准样","50","51.174","2.35%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","化学需氧量","2023-05-11","实际样","25","24.79","-0.21mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","化学需氧量","2023-05-11","实际样","25","23.28","-1.72mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","化学需氧量","2023-05-11","实际样","25","22.65","-2.35mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","化学需氧量","2023-05-11","标准样","50","53.85","7.7%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","氨氮","2023-05-11","实际样","1.5","1.382","-0.12mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","氨氮","2023-05-11","实际样","1.5","1.550","0.05mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","氨氮","2023-05-11","实际样","1.5","1.306","-0.19mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","氨氮","2023-05-11","标准样","5","5.068","1.36%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总磷","2023-05-11","实际样","0.2","0.224","0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总磷","2023-05-11","实际样","0.2","0.214","0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总磷","2023-05-11","实际样","0.2","0.225","0.02mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总磷","2023-05-11","标准样","0.5","0.489","-2.2%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总氮","2023-05-11","实际样","5.61","5.79","3.21%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总氮","2023-05-11","实际样","5.46","5.79","6.04%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总氮","2023-05-11","实际样","5.42","5.57","2.77%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","总氮","2023-05-11","标准样","10","10.08","0.8%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-012苏州市相润排水管理有限公司（一泓污水处理厂）","苏州市相润排水管理有限公司一泓污水处理厂","废水总排口(S1)","pH值","2023-05-11","实际样","7.3","6.83","-0.47pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","氨氮","2023-05-09","实际样","1.5","1.494","-0.01mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","氨氮","2023-05-09","实际样","1.5","1.484","-0.02mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","氨氮","2023-05-09","实际样","1.5","1.490","-0.01mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","氨氮","2023-05-09","标准样","5","5.393","7.86%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","化学需氧量","2023-05-09","实际样","25","23.78","-1.22mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","化学需氧量","2023-05-09","实际样","25","23.35","-1.65mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","化学需氧量","2023-05-09","实际样","25","23.18","-1.82mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","化学需氧量","2023-05-09","标准样","50","50.09","0.18%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总磷","2023-05-09","实际样","0.2","0.205","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总磷","2023-05-09","实际样","0.2","0.202","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总磷","2023-05-09","实际样","0.2","0.203","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总磷","2023-05-09","标准样","0.5","0.508","1.6%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总氮","2023-05-09","实际样","5.79","4.96","-14.34%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总氮","2023-05-09","实际样","5.48","5.77","5.29%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总氮","2023-05-09","实际样","5.70","4.96","-12.98%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","总氮","2023-05-09","标准样","10","9.51","-4.9%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-011苏州市相润排水管理有限公司（漕湖污水处理厂）","苏州市相润排水管理有限公司（漕湖污水处理厂）","废水总排口(S1)","pH值","2023-05-09","实际样","7.0","6.72","-0.28pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-021苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-02-07","实际样","0.2","0.215","0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-021苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-02-07","实际样","0.2","0.209","0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-021苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-02-07","实际样","0.2","0.207","0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-021苏州高铁苏水水务有限公司","苏州高铁苏水水务有限公司","高铁新城污水排放口","总磷（以P计）","2023-02-07","标准样","0.5","0.510","2%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总铬","2023-02-02","实际样","<0.03","0.015","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总铬","2023-02-02","实际样","<0.03","0.014","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","总铬","2023-02-02","实际样","<0.03","0.013","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","铜","2023-02-02","实际样","<0.04","0.008","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","铜","2023-02-02","实际样","<0.04","0.006","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","铜","2023-02-02","实际样","<0.04","0.008","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","镍 ","2023-02-02","实际样","<0.007","0.037","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","镍 ","2023-02-02","实际样","<0.007","0.046","",""]
["重点源企业上报（比对）2023-020苏州市相润排水管理有限公司黄桥分公司","苏州市相润排水管理有限公司黄桥分公司","废水排放口S1(S1)","镍 ","2023-02-02","实际样","<0.007","0.050","",""]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","氨氮","2023-02-01","实际样","1.5","1.443","-0.06mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","氨氮","2023-02-01","实际样","1.5","1.516","0.02mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","氨氮","2023-02-01","实际样","1.5","1.525","0.02mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","氨氮","2023-02-01","标准样","5","4.942","-1.16%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总氮","2023-02-01","实际样","8.72","8.442","-3.19%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总氮","2023-02-01","实际样","8.61","8.554","-0.65%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总氮","2023-02-01","实际样","8.46","8.176","-3.36%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总氮","2023-02-01","标准样","15","15.552","3.68%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","化学需氧量","2023-02-01","实际样","25","26.9","1.9mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","化学需氧量","2023-02-01","实际样","25","27.5","2.5mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","化学需氧量","2023-02-01","实际样","25","24.3","-0.7mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","化学需氧量","2023-02-01","标准样","50","47.5","-5%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总磷","2023-02-01","实际样","0.2","0.200","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总磷","2023-02-01","实际样","0.2","0.201","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总磷","2023-02-01","实际样","0.2","0.200","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","总磷","2023-02-01","标准样","0.5","0.503","0.6%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-019苏州市相润排水管理有限公司（澄阳污水处理厂）","苏州市相润排水管理有限公司澄阳污水处理厂","总排口(S1)","pH值","2023-02-01","实际样","7.2","6.99","-0.21pH","绝对误差不超过±0.5pH"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总磷","2023-01-11","实际样","0.2","0.191","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总磷","2023-01-11","实际样","0.2","0.202","0mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总磷","2023-01-11","实际样","0.2","0.195","-0.01mg/L","绝对误差不超过±0.04mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总磷","2023-01-11","标准样","0.5","0.504","0.8%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","氨氮","2023-01-11","实际样","1.5","1.506","0.01mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","氨氮","2023-01-11","实际样","1.5","1.464","-0.04mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","氨氮","2023-01-11","实际样","1.5","1.535","0.03mg/L","绝对误差不超过±0.3mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","氨氮","2023-01-11","标准样","5","5.129","2.58%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","化学需氧量","2023-01-11","实际样","25","28.1","3.1mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","化学需氧量","2023-01-11","实际样","25","27.0","2mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","化学需氧量","2023-01-11","实际样","25","28.0","3mg/L","绝对误差不超过±5mg/L"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","化学需氧量","2023-01-11","标准样","50","51.6","3.2%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总氮","2023-01-11","实际样","7.05","6.56","-6.95%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总氮","2023-01-11","实际样","7.01","6.557","-6.46%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总氮","2023-01-11","实际样","6.58","7.35","11.7%","相对误差不超过±15%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","总氮","2023-01-11","标准样","15","14.51","-3.27%","相对误差不超过±10%"]
["重点源企业上报（比对）2023-018苏州市相城区东桥集中污水处理厂","苏州市相城区东桥集中污水处理厂","污水排放口(S1)","pH值","2023-01-11","实际样","7.5","7.38","-0.12pH","绝对误差不超过±0.5pH"]